## Header
| Version | V2.2
| Published | 2021-04-28
Disclaimer - ams AG
This Excel sheet is a calibration aid that shows the calibration based on an XYZ calibration. Furthermore, the file shows the path in the calculation "from ADC value to calculating xyz, ab, and Lab (D65 / 2° observer), based on 24 colors of the ColorChecker". All the data in Excel are examples and were measured based on an AS7341 test board. Therefore they have no general validity for all sensor boards or only reduced accuracy for sensor boards of identical construction. Only use the original downloaded versions of this file and do not send, save, use, or manipulate a copy of it. Versions of this document with a different author or publication date are not the original file from ams and hence are not supported. All the tables are not read-only and therefore not tamper-proof. Thus every detail is only valid as information if it comes directly from the ams download and has not been modified.
IMPORTANT - PLEASE READ CAREFULLY BEFORE COPYING, INSTALLING, OR USING THE SOFTWARE. THIS SOFTWARE IS PROVIDED BY THE COPYRIGHT HOLDERS (ams AG – all rights reserved) AND CONTRIBUTORS "AS IS" AND ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, BUT NOT LIMITED TO, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE ARE DISCLAIMED. IN NO EVENT SHALL THE COPYRIGHT OWNER OR CONTRIBUTORS BE LIABLE FOR ANY DIRECT, INDIRECT, INCIDENTAL, SPECIAL, EXEMPLARY, OR CONSEQUENTIAL DAMAGES (INCLUDING, BUT NOT LIMITED TO PROCUREMENT OF SUBSTITUTE GOODS OR SERVICES; LOSS OF USE, DATA, OR PROFITS; OR BUSINESS INTERRUPTION) HOWEVER CAUSED AND ON ANY THEORY OF LIABILITY, WHETHER IN CONTRACT, STRICT LIABILITY, OR TORT (INCLUDING NEGLIGENCE OR OTHERWISE) ARISING IN ANY WAY OUT OF THE USE OF THIS SOFTWARE, EVEN IF ADVISED OF THE POSSIBILITY OF SUCH DAMAGE.
I) See the formulas in the sheet 'CM calculation' for the Calculation of a Correction Matrix based on the Spectrometer and Sensor results.
II) Copy your measured results into the sheet 'CM Calculation' if you use the excel sheets to generate your calibration matrix.
III) See the test results for 24 colored targets, based on the actual calculated Correction Matrix, in the sheet 'Test Results', from line 74 and its following.
IV) See all used formulas from ADC to XYZ and Lab in the sheet 'Test Results'.

## CM Calculation
Calculation 'Correction Matrix' |  | Yellow cells = input cells
Action 1: Copy the Spectrometer XYZ values of the measured reference colors into cells [c6:z8] OR use the values here as reference values AND accept a lower accuracy depending on the differences between your ColorChecker and the values used here.
TARGET DATA T (Reference Values) | Reference values | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
|  | Dark Skin | Light Skin | Blue Sky | Foliage | Blue Flower | Bluish Green | Orange Yellow | Yellow Green | Purple | Moderate Red | Purplish Blue | Orange | Blue | Green | Red | Yellow | Magenta | Cyan | Black | Neutral 3.5 | Neutral 5 | Neutral 6.5 | Neutral 8 | White
| X | 11.22 | 36.35 | 16.9 | 10.93 | 24.56 | 30.76 | 46.31 | 33.91 | 8.51 | 28.12 | 13.93 | 37.82 | 8.44 | 14.23 | 19.98 | 55.79 | 29.69 | 14.5 | 3.08 | 8.69 | 18.35 | 35.15 | 56.93 | 83.83
| Y | 10.23 | 33.51 | 17.92 | 13.63 | 22.76 | 42.37 | 43.45 | 44.63 | 6.39 | 18.87 | 11.89 | 30.31 | 6.38 | 22.91 | 12.2 | 59.75 | 19.31 | 19.66 | 3.25 | 9.2 | 19.34 | 37.12 | 60 | 88.53
| Z | 7.15 | 24.55 | 33.54 | 7.13 | 43.25 | 45.05 | 7.83 | 11.31 | 14.41 | 13.56 | 39.96 | 6.28 | 29.75 | 9.82 | 5.02 | 9.03 | 31.05 | 39.76 | 3.68 | 10.23 | 20.82 | 40.19 | 64.74 | 92.21
Offset Sensor (e.g. Optical Overcrossing) | Offset | 1: Sensor; Action 1: Copy the Spectrometer XYZ values of the measured reference colors into cells [c6:z8] OR use the values here as reference values AND accept a lower accuracy depending on the differences between your ColorChecker and the values used here. 2: Action 2: Copy the 'Basic Counts' of Offset into the Cells [C11:C18]. In our example, the values represent the direct radiation of the used LEDs into the sensor, measured in a DARK room.
| 415 | 1: 0.00372
| 445 | 1: 0.013924; Action 1: Copy the Spectrometer XYZ values of the measured reference colors into cells [c6:z8] OR use the values here as reference values AND accept a lower accuracy depending on the differences between your ColorChecker and the values used here. 3: Background of the calibration is a calculation of target values T and sensor values S with the method of linear regression. The target values are in XYZ, measured with a spectrometer or known standard values of colors, which are measured with a sensor as sensor raw data. The sensor raw data are corrected (only offset is used here) and then referenced with the spectrometer data. The result of the calibration is a correction matrix CM (see [c55:j57]), which can be used in the future under the same conditions as in the calibration matching of the sensor values. Realize action 1...4 in the calculation of a customized calibration matrix or follow the formulas backwards from calibration matrix to input values to understand the calculations.
| 480 | 1: 0.0064408
| 515 | 1: 0.028512
| 555 | 1: 0.030216
| 590 | 1: 0.030085
| 630 | 1: 0.040693
| 680 | 1: 0.012133
Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30].
Sensor Matrix S | Basic Counts | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
|  | Dark Skin | Light Skin | Blue Sky | Foliage | Blue Flower | Bluish Green | Orange Yellow | Yellow Green | Purple | Moderate Red | Purplish Blue | Orange | Blue | Green | Red | Yellow | Magenta | Cyan | Black | Neutral 3.5 | Neutral 5 | Neutral 6.5 | Neutral 8 | White
| 415 | 0.029475 | 0.087527 | 0.077362 | 0.027406 | 0.10674 | 0.10792 | 0.074949 | 0.066278 | 0.040942 | 0.058922 | 0.081729 | 0.058337 | 0.058721 | 0.036519 | 0.035374 | 0.091089 | 0.096882 | 0.073214 | 0.013401 | 0.029787 | 0.055945 | 0.10862 | 0.17832 | 0.25576
| 445 | 0.086782 | 0.25168 | 0.32701 | 0.075782 | 0.42676 | 0.39957 | 0.12075 | 0.1229 | 0.14801 | 0.15926 | 0.38811 | 0.097485 | 0.30204 | 0.091774 | 0.073546 | 0.13067 | 0.32839 | 0.34923 | 0.048459 | 0.10746 | 0.19964 | 0.3857 | 0.63196 | 0.90192
| 480 | 0.049863 | 0.18179 | 0.1306 | 0.053574 | 0.16627 | 0.24859 | 0.11736 | 0.14084 | 0.047627 | 0.087662 | 0.11418 | 0.092678 | 0.082321 | 0.095101 | 0.053453 | 0.15389 | 0.11721 | 0.17828 | 0.023039 | 0.051277 | 0.096658 | 0.1876 | 0.30744 | 0.45141
| 515 | 0.18989 | 0.63372 | 0.42114 | 0.30315 | 0.51243 | 1.0655 | 0.63308 | 1.0068 | 0.13481 | 0.24007 | 0.26595 | 0.32877 | 0.16116 | 0.62109 | 0.14665 | 1.0711 | 0.26599 | 0.56784 | 0.091772 | 0.19972 | 0.37407 | 0.72533 | 1.192 | 1.7504
| 555 | 0.31438 | 0.85895 | 0.46402 | 0.39605 | 0.49125 | 1.1161 | 1.3798 | 1.3968 | 0.14992 | 0.37833 | 0.26163 | 0.91308 | 0.15961 | 0.65867 | 0.2182 | 1.9083 | 0.37197 | 0.39621 | 0.11715 | 0.26681 | 0.51809 | 0.99577 | 1.6353 | 2.4182
| 590 | 0.40503 | 1.3053 | 0.41656 | 0.32049 | 0.56569 | 0.74547 | 1.6949 | 1.0633 | 0.21544 | 1.0312 | 0.24935 | 1.5104 | 0.15926 | 0.38168 | 0.65727 | 1.9729 | 0.91244 | 0.27396 | 0.11604 | 0.25997 | 0.50941 | 0.97513 | 1.6073 | 2.3925
| 630 | 0.30905 | 1.0479 | 0.29192 | 0.22727 | 0.63343 | 0.38996 | 1.1712 | 0.61093 | 0.26449 | 1.0174 | 0.21101 | 1.0424 | 0.13896 | 0.21174 | 0.96589 | 1.312 | 1.0188 | 0.19996 | 0.097833 | 0.18809 | 0.34652 | 0.65 | 1.0701 | 1.601
| 680 | 0.10734 | 0.40668 | 0.11507 | 0.081833 | 0.26913 | 0.15322 | 0.41703 | 0.23949 | 0.10769 | 0.34687 | 0.12062 | 0.35582 | 0.060977 | 0.084673 | 0.35294 | 0.46807 | 0.41042 | 0.079972 | 0.032698 | 0.064293 | 0.11906 | 0.23183 | 0.38951 | 0.58886
Sensor Matrix corr offset S' = S - offset | Basic-Offset | Dark Skin | Light Skin | Blue Sky | Foliage | Blue Flower | Bluish Green | Orange Yellow | Yellow Green | Purple | Moderate Red | Purplish Blue | Orange | Blue | Green | Red | Yellow | Magenta | Cyan | Black | Neutral 3.5 | Neutral 5 | Neutral 6.5 | Neutral 8 | White
| 415 | 0.025755 | 0.083807 | 0.073642 | 0.023686 | 0.10302 | 0.1042 | 0.071229 | 0.062558 | 0.037222 | 0.055202 | 0.078009 | 0.054617 | 0.055001 | 0.032799 | 0.031654 | 0.087369 | 0.093162 | 0.069494 | 0.009681 | 0.026067 | 0.052225 | 0.1049 | 0.1746 | 0.25204
| 445 | 0.072858 | 0.23775 | 0.31308 | 0.061858 | 0.41283 | 0.38564 | 0.10683 | 0.10897 | 0.13408 | 0.14533 | 0.37419 | 0.083561 | 0.28811 | 0.07785 | 0.059622 | 0.11674 | 0.31447 | 0.3353 | 0.034535 | 0.093537 | 0.18571 | 0.37178 | 0.61804 | 0.888
| 480 | 0.043422 | 0.17535 | 0.12416 | 0.047133 | 0.15983 | 0.24215 | 0.11092 | 0.1344 | 0.041186 | 0.081221 | 0.10774 | 0.086237 | 0.07588 | 0.08866 | 0.047012 | 0.14745 | 0.11077 | 0.17184 | 0.016598 | 0.044836 | 0.090217 | 0.18115 | 0.301 | 0.44497
| 515 | 0.16137 | 0.60521 | 0.39263 | 0.27464 | 0.48391 | 1.037 | 0.60456 | 0.97828 | 0.1063 | 0.21156 | 0.23744 | 0.30026 | 0.13265 | 0.59258 | 0.11814 | 1.0426 | 0.23748 | 0.53933 | 0.06326 | 0.17121 | 0.34556 | 0.69682 | 1.1635 | 1.7219
| 555 | 0.28416 | 0.82874 | 0.4338 | 0.36584 | 0.46104 | 1.0859 | 1.3496 | 1.3666 | 0.1197 | 0.34811 | 0.23142 | 0.88287 | 0.12939 | 0.62846 | 0.18798 | 1.8781 | 0.34175 | 0.36599 | 0.086934 | 0.23659 | 0.48787 | 0.96556 | 1.6051 | 2.3879
| 590 | 0.37494 | 1.2752 | 0.38647 | 0.29041 | 0.53561 | 0.71538 | 1.6649 | 1.0332 | 0.18536 | 1.0011 | 0.21926 | 1.4803 | 0.12917 | 0.35159 | 0.62719 | 1.9428 | 0.88235 | 0.24388 | 0.085953 | 0.22989 | 0.47932 | 0.94505 | 1.5772 | 2.3624
| 630 | 0.26836 | 1.0072 | 0.25123 | 0.18658 | 0.59274 | 0.34927 | 1.1305 | 0.57023 | 0.2238 | 0.97668 | 0.17032 | 1.0017 | 0.09827 | 0.17104 | 0.9252 | 1.2713 | 0.97806 | 0.15927 | 0.05714 | 0.14739 | 0.30583 | 0.60931 | 1.0294 | 1.5603
| 680 | 0.095203 | 0.39455 | 0.10294 | 0.0697 | 0.257 | 0.14108 | 0.4049 | 0.22735 | 0.095556 | 0.33474 | 0.10849 | 0.34368 | 0.048844 | 0.07254 | 0.3408 | 0.45593 | 0.39828 | 0.067839 | 0.020565 | 0.05216 | 0.10693 | 0.2197 | 0.37738 | 0.57673
Calculation: Matrix for linear transformation CM
| A = T * S'_trans | 1: 67.1967; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 2: 212.1345; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 3: 117.6892; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 4: 487.9744; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 5: 716.4277; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 6: 785.1816; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 7: 554.4346; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 8: 206.2904
|  | 69.0612 | 219.0383 | 123.399 | 524.1187 | 755.3897 | 790.9184 | 542.4265 | 202.059
|  | 68.7518 | 249.5483 | 121.031 | 450.2365 | 584.5516 | 599.0902 | 423.1566 | 162.2731
| B = [S' * S'_trans]^-1 | 1: 9143.9372; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 2: -1887.9981; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 3: 772.0494; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 4: -177.7958; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 5: -85.8054; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 6: -88.149; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 7: -51.8521; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 8: -283.6122
|  | -1887.9981 | 434.8988 | -356.178 | 82.4778 | -11.0488 | 41.494 | 14.7962 | 16.5615
|  | 772.0494 | -356.178 | 994.7261 | -248.8236 | 143.9939 | -128.2022 | -2.7683 | 131.1707
|  | -177.7958 | 82.4778 | -248.8236 | 71.2653 | -45.7156 | 38.7455 | -2.8455 | -34.572
|  | -85.8054 | -11.0488 | 143.9939 | -45.7156 | 37.2976 | -29.608 | 9.883 | 21.9438
|  | -88.149 | 41.494 | -128.2022 | 38.7455 | -29.608 | 29.496 | -15.955 | 1.0179
|  | -51.8521 | 14.7962 | -2.7683 | -2.8455 | 9.883 | -15.955 | 50.0347 | -98.0004
|  | -283.6122 | 16.5615 | 131.1707 | -34.572 | 21.9438 | 1.0179 | -98.0004 | 265.5054
Action 4: See and use the calculated Correction Matrix.
Calculated Correction Matrix | CM = A * B | 1: 93.3544; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 2: 3.2562; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 3: -5.8807; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 4: 1.6336; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 5: 8.4827; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 6: 9.4743; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 7: 17.4032; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 8: -20.6196; Action 3: Copy the Sensor 'Basic Counts' from the protocol log file of all 24 targets into the Cells [C23:Z30]. 10: Copy this matrix CM into the CM - File of AS7341 EVK Reflection Mode when using a customized calibration.
|  | 64.0772 | -7.3362 | 13.8011 | 16.2001 | 13.8366 | 2.9134 | 11.605 | -21.5623
|  | -23.483 | 99.493 | 40.6773 | 3.939 | -4.1265 | -0.39083 | -0.48666 | -3.863

## Test Results
Yellow cells = input cells | Calculation 'Test Results' |  |  |  | Action 1: Copy the Spectrometer XYZ values of the spectrometer measured colors into cells [c5:z7] OR use the values from the calibration sheet to verify the calibrated colors.
Spectrometer | Reference values | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
| Spectrometer | Dark Skin | Light Skin | Blue Sky | Foliage | Blue Flower | Bluish Green | Orange Yellow | Yellow Green | Purple | Moderate Red | Purplish Blue | Orange | Blue | Green | Red | Yellow | Magenta | Cyan | Black | Neutral 3.5 | Neutral 5 | Neutral 6.5 | Neutral 8 | White
| X | 11.22 | 36.35 | 16.9 | 10.93 | 24.56 | 30.76 | 46.31 | 33.91 | 8.51 | 28.12 | 13.93 | 37.82 | 8.44 | 14.23 | 19.98 | 55.79 | 29.69 | 14.5 | 3.08 | 8.69 | 18.35 | 35.15 | 56.93 | 83.83
| Y | 10.23 | 33.51 | 17.92 | 13.63 | 22.76 | 42.37 | 43.45 | 44.63 | 6.39 | 18.87 | 11.89 | 30.31 | 6.38 | 22.91 | 12.2 | 59.75 | 19.31 | 19.66 | 3.25 | 9.2 | 19.34 | 37.12 | 60 | 88.53
| Z | 7.15 | 24.55 | 33.54 | 7.13 | 43.25 | 45.05 | 7.83 | 11.31 | 14.41 | 13.56 | 39.96 | 6.28 | 29.75 | 9.82 | 5.02 | 9.03 | 31.05 | 39.76 | 3.68 | 10.23 | 20.82 | 40.19 | 64.74 | 92.21
| Reference values | Reference values Dark Skin | Reference values Light Skin | Reference values Blue Sky | Reference values Foliage | Reference values Blue Flower | Reference values Bluish Green | Reference values Orange Yellow | Reference values Yellow Green | Reference values Purple | Reference values Moderate Red | Reference values Purplish Blue | Reference values Orange | Reference values Blue | Reference values Green | Reference values Red | Reference values Yellow | Reference values Magenta | Reference values Cyan | Reference values Black | Reference values Neutral 3.5 | Reference values Neutral 5 | Reference values Neutral 6.5 | Reference values Neutral 8 | Reference values White
| x | 0.39231 | 0.38502 | 0.24722 | 0.3449 | 0.27117 | 0.26028 | 0.47454 | 0.37741 | 0.29034 | 0.46441 | 0.21177 | 0.50827 | 0.18937 | 0.30302 | 0.5371 | 0.44786 | 0.37089 | 0.19616 | 0.30769 | 0.30903 | 0.31362 | 0.31256 | 0.31337 | 0.31685
| y | 0.35769 | 0.35494 | 0.26214 | 0.4301 | 0.2513 | 0.35852 | 0.44523 | 0.49672 | 0.21801 | 0.31164 | 0.18075 | 0.40734 | 0.14315 | 0.48786 | 0.32796 | 0.47965 | 0.24122 | 0.26596 | 0.32468 | 0.32717 | 0.33054 | 0.33007 | 0.33027 | 0.33462
| z | 0.25 | 0.26004 | 0.49064 | 0.22499 | 0.47753 | 0.3812 | 0.080234 | 0.12588 | 0.49164 | 0.22395 | 0.60748 | 0.084397 | 0.66749 | 0.20911 | 0.13495 | 0.072489 | 0.38788 | 0.53788 | 0.36763 | 0.3638 | 0.35584 | 0.35737 | 0.35636 | 0.34853
Sensor Values and Results | Sensor Values | Action 1: Copy the Spectrometer XYZ values of the spectrometer measured colors into cells [c5:z7] OR use the values from the calibration sheet to verify the calibrated colors. 4: Action 2: Copy the 'Basic Counts' of Offset into the Cells [C16:C23] or use the offset from the calibration sheet.
| Offset | 1: Dark Basic Counts
| 415 | 1: 0.00372; Action 1: Copy the Spectrometer XYZ values of the spectrometer measured colors into cells [c5:z7] OR use the values from the calibration sheet to verify the calibrated colors. 8: The correction matrix is used to match the immediate sensor results in the application - here the color space is used to define target colors and to check the accuracy with reference values. This can be done with the colors and values from the calibration or with other colors, each measured with spectrometer and sensor (follow action 1...3). The data chain in the sheet is that the sensor values multiplied by the matrix results in the sensor XYZ values. These are compared with the spectrometer values as XYZ or further down as xy and Lab. The errors are output as Delta per color or Max/Min/Average.
| 445 | 1: 0.013924
| 480 | 1: 0.0064408
| 515 | 1: 0.028512
| 555 | 1: 0.030216
| 590 | 1: 0.030085
| 630 | 1: 0.040693
| 680 | 1: 0.012133
Action 3: Copy the Sensor 'Basic Counts' from the protocol log file or use them from the calibration sheet to verify the calibrated colors.
| Sensor Values | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
| Basic counts | Dark Skin | Light Skin | Blue Sky | Foliage | Blue Flower | Bluish Green | Orange Yellow | Yellow Green | Purple | Moderate Red | Purplish Blue | Orange | Blue | Green | Red | Yellow | Magenta | Cyan | Black | Neutral 3.5 | Neutral 5 | Neutral 6.5 | Neutral 8 | White
| 415 | 0.029475 | 0.087527 | 0.077362 | 0.027406 | 0.10674 | 0.10792 | 0.074949 | 0.066278 | 0.040942 | 0.058922 | 0.081729 | 0.058337 | 0.058721 | 0.036519 | 0.035374 | 0.091089 | 0.096882 | 0.073214 | 0.013401 | 0.029787 | 0.055945 | 0.10862 | 0.17832 | 0.25576
| 445 | 0.086782 | 0.25168 | 0.32701 | 0.075782 | 0.42676 | 0.39957 | 0.12075 | 0.1229 | 0.14801 | 0.15926 | 0.38811 | 0.097485 | 0.30204 | 0.091774 | 0.073546 | 0.13067 | 0.32839 | 0.34923 | 0.048459 | 0.10746 | 0.19964 | 0.3857 | 0.63196 | 0.90192
| 480 | 0.049863 | 0.18179 | 0.1306 | 0.053574 | 0.16627 | 0.24859 | 0.11736 | 0.14084 | 0.047627 | 0.087662 | 0.11418 | 0.092678 | 0.082321 | 0.095101 | 0.053453 | 0.15389 | 0.11721 | 0.17828 | 0.023039 | 0.051277 | 0.096658 | 0.1876 | 0.30744 | 0.45141
| 515 | 0.18989 | 0.63372 | 0.42114 | 0.30315 | 0.51243 | 1.0655 | 0.63308 | 1.0068 | 0.13481 | 0.24007 | 0.26595 | 0.32877 | 0.16116 | 0.62109 | 0.14665 | 1.0711 | 0.26599 | 0.56784 | 0.091772 | 0.19972 | 0.37407 | 0.72533 | 1.192 | 1.7504
| 555 | 0.31438 | 0.85895 | 0.46402 | 0.39605 | 0.49125 | 1.1161 | 1.3798 | 1.3968 | 0.14992 | 0.37833 | 0.26163 | 0.91308 | 0.15961 | 0.65867 | 0.2182 | 1.9083 | 0.37197 | 0.39621 | 0.11715 | 0.26681 | 0.51809 | 0.99577 | 1.6353 | 2.4182
| 590 | 0.40503 | 1.3053 | 0.41656 | 0.32049 | 0.56569 | 0.74547 | 1.6949 | 1.0633 | 0.21544 | 1.0312 | 0.24935 | 1.5104 | 0.15926 | 0.38168 | 0.65727 | 1.9729 | 0.91244 | 0.27396 | 0.11604 | 0.25997 | 0.50941 | 0.97513 | 1.6073 | 2.3925
| 630 | 0.30905 | 1.0479 | 0.29192 | 0.22727 | 0.63343 | 0.38996 | 1.1712 | 0.61093 | 0.26449 | 1.0174 | 0.21101 | 1.0424 | 0.13896 | 0.21174 | 0.96589 | 1.312 | 1.0188 | 0.19996 | 0.097833 | 0.18809 | 0.34652 | 0.65 | 1.0701 | 1.601
| 680 | 0.10734 | 0.40668 | 0.11507 | 0.081833 | 0.26913 | 0.15322 | 0.41703 | 0.23949 | 0.10769 | 0.34687 | 0.12062 | 0.35582 | 0.060977 | 0.084673 | 0.35294 | 0.46807 | 0.41042 | 0.079972 | 0.032698 | 0.064293 | 0.11906 | 0.23183 | 0.38951 | 0.58886
| Corrected counts | Dark Skin | Light Skin | Blue Sky | Foliage | Blue Flower | Bluish Green | Orange Yellow | Yellow Green | Purple | Moderate Red | Purplish Blue | Orange | Blue | Green | Red | Yellow | Magenta | Cyan | Black | Neutral 3.5 | Neutral 5 | Neutral 6.5 | Neutral 8 | White
| 415 | 0.025755 | 0.083807 | 0.073642 | 0.023686 | 0.10302 | 0.1042 | 0.071229 | 0.062558 | 0.037222 | 0.055202 | 0.078009 | 0.054617 | 0.055001 | 0.032799 | 0.031654 | 0.087369 | 0.093162 | 0.069494 | 0.009681 | 0.026067 | 0.052225 | 0.1049 | 0.1746 | 0.25204
| 445 | 0.072858 | 0.23775 | 0.31308 | 0.061858 | 0.41283 | 0.38564 | 0.10683 | 0.10897 | 0.13408 | 0.14533 | 0.37419 | 0.083561 | 0.28811 | 0.07785 | 0.059622 | 0.11674 | 0.31447 | 0.3353 | 0.034535 | 0.093537 | 0.18571 | 0.37178 | 0.61804 | 0.888
| 480 | 0.043422 | 0.17535 | 0.12416 | 0.047133 | 0.15983 | 0.24215 | 0.11092 | 0.1344 | 0.041186 | 0.081221 | 0.10774 | 0.086237 | 0.07588 | 0.08866 | 0.047012 | 0.14745 | 0.11077 | 0.17184 | 0.016598 | 0.044836 | 0.090217 | 0.18115 | 0.301 | 0.44497
| 515 | 0.16137 | 0.60521 | 0.39263 | 0.27464 | 0.48391 | 1.037 | 0.60456 | 0.97828 | 0.1063 | 0.21156 | 0.23744 | 0.30026 | 0.13265 | 0.59258 | 0.11814 | 1.0426 | 0.23748 | 0.53933 | 0.06326 | 0.17121 | 0.34556 | 0.69682 | 1.1635 | 1.7219
| 555 | 0.28416 | 0.82874 | 0.4338 | 0.36584 | 0.46104 | 1.0859 | 1.3496 | 1.3666 | 0.1197 | 0.34811 | 0.23142 | 0.88287 | 0.12939 | 0.62846 | 0.18798 | 1.8781 | 0.34175 | 0.36599 | 0.086934 | 0.23659 | 0.48787 | 0.96556 | 1.6051 | 2.3879
| 590 | 0.37494 | 1.2752 | 0.38647 | 0.29041 | 0.53561 | 0.71538 | 1.6649 | 1.0332 | 0.18536 | 1.0011 | 0.21926 | 1.4803 | 0.12917 | 0.35159 | 0.62719 | 1.9428 | 0.88235 | 0.24388 | 0.085953 | 0.22989 | 0.47932 | 0.94505 | 1.5772 | 2.3624
| 630 | 0.26836 | 1.0072 | 0.25123 | 0.18658 | 0.59274 | 0.34927 | 1.1305 | 0.57023 | 0.2238 | 0.97668 | 0.17032 | 1.0017 | 0.09827 | 0.17104 | 0.9252 | 1.2713 | 0.97806 | 0.15927 | 0.05714 | 0.14739 | 0.30583 | 0.60931 | 1.0294 | 1.5603
| 680 | 0.095203 | 0.39455 | 0.10294 | 0.0697 | 0.257 | 0.14108 | 0.4049 | 0.22735 | 0.095556 | 0.33474 | 0.10849 | 0.34368 | 0.048844 | 0.07254 | 0.3408 | 0.45593 | 0.39828 | 0.067839 | 0.020565 | 0.05216 | 0.10693 | 0.2197 | 0.37738 | 0.57673
| Calibrated Results | Dark Skin | Light Skin | Blue Sky | Foliage | Blue Flower | Bluish Green | Orange Yellow | Yellow Green | Purple | Moderate Red | Purplish Blue | Orange | Blue | Green | Red | Yellow | Magenta | Cyan | Black | Neutral 3.5 | Neutral 5 | Neutral 6.5 | Neutral 8 | White
| X | 11.3199 | 37.0598 | 17.3964 | 10.2486 | 24.8136 | 30.4117 | 45.8807 | 33.6199 | 8.5389 | 28.027 | 13.0227 | 37.2149 | 8.8677 | 13.9051 | 19.6767 | 56.4335 | 29.5252 | 14.238 | 3.1441 | 8.4286 | 17.3114 | 34.2948 | 57.1337 | 84.5162
| Y | 10.4151 | 34.2135 | 18.3204 | 12.7339 | 22.8946 | 42.1094 | 43.0192 | 44.5468 | 6.4251 | 18.8691 | 11.065 | 29.6843 | 6.8608 | 22.4949 | 11.9703 | 60.236 | 19.1003 | 19.2621 | 3.2938 | 8.9057 | 18.2177 | 36.2304 | 60.2696 | 89.267
| Z | 7.2285 | 25.271 | 33.5565 | 6.6141 | 43.6696 | 44.381 | 7.515 | 11.4947 | 13.5159 | 13.7044 | 39.172 | 5.6849 | 30.1617 | 9.8217 | 4.7787 | 8.7791 | 30.7712 | 38.8976 | 3.6334 | 9.8529 | 19.5193 | 39.1406 | 65.0184 | 93.5497
| x | 0.39083 | 0.38386 | 0.25113 | 0.34628 | 0.27155 | 0.26015 | 0.47587 | 0.37496 | 0.29982 | 0.46249 | 0.20586 | 0.51271 | 0.19324 | 0.30083 | 0.54019 | 0.44985 | 0.37187 | 0.19666 | 0.31218 | 0.31002 | 0.31448 | 0.31272 | 0.3132 | 0.31615
| y | 0.35959 | 0.35438 | 0.26447 | 0.43025 | 0.25055 | 0.36021 | 0.44619 | 0.49683 | 0.2256 | 0.31137 | 0.17491 | 0.40896 | 0.1495 | 0.48667 | 0.32862 | 0.48016 | 0.24057 | 0.26606 | 0.32705 | 0.32757 | 0.33094 | 0.33037 | 0.33039 | 0.33392
| z | 0.24957 | 0.26176 | 0.48441 | 0.22347 | 0.4779 | 0.37964 | 0.077945 | 0.1282 | 0.47458 | 0.22614 | 0.61923 | 0.078322 | 0.65726 | 0.21249 | 0.13119 | 0.069982 | 0.38756 | 0.53728 | 0.36077 | 0.36241 | 0.35458 | 0.35691 | 0.35642 | 0.34994
Diff XYZ | XSP-XS | -0.8901% | -1.953% | -2.938% | 6.234% | -1.032% | 1.132% | 0.9269% | 0.8555% | -0.3401% | 0.3308% | 6.513% | 1.6% | -5.068% | 2.283% | 1.518% | -1.154% | 0.5552% | 1.807% | -2.081% | 3.008% | 5.66% | 2.433% | -0.3579% | -0.8186%
| YSP-XS | -1.809% | -2.099% | -2.234% | 6.574% | -0.5915% | 0.615% | 0.9914% | 0.1863% | -0.5488% | 0.004964% | 6.939% | 2.064% | -7.536% | 1.812% | 1.883% | -0.8134% | 1.086% | 2.024% | -1.349% | 3.198% | 5.803% | 2.397% | -0.4493% | -0.8325%
| ZSP-XS | -1.098% | -2.937% | -0.04908% | 7.236% | -0.9701% | 1.485% | 4.023% | -1.633% | 6.205% | -1.065% | 1.972% | 9.476% | -1.384% | -0.01779% | 4.808% | 2.778% | 0.898% | 2.169% | 1.266% | 3.686% | 6.247% | 2.611% | -0.43% | -1.453%
Diff xy | dxy | 0.0024061 | 0.0012878 | 0.0045459 | 0.001381 | 0.00083801 | 0.0016953 | 0.00164 | 0.0024448 | 0.01214 | 0.0019413 | 0.0083061 | 0.0047373 | 0.0074451 | 0.0024912 | 0.0031617 | 0.0020584 | 0.0011756 | 0.00051472 | 0.0050805 | 0.001066 | 0.00094244 | 0.00034068 | 0.00020987 | 0.00099655
| Average dxy | 1: 0.0028686
| Min dxy | 1: 0.00020987
| Max dxy | 1: 0.01214
Lab results | Spektrometer
| L (D65/2°) | 38.2521 | 64.5718 | 49.3989 | 43.6976 | 54.8243 | 71.1253 | 71.8593 | 72.6476 | 30.3758 | 50.5348 | 41.0408 | 61.9208 | 30.3516 | 54.9795 | 41.5322 | 81.7021 | 51.0479 | 51.4506 | 21.0185 | 36.3665 | 51.0826 | 67.367 | 81.8382 | 95.3837
| a (D65/2°) | 11.4306 | 15.64 | -0.73245 | -14.1729 | 13.1916 | -32.2567 | 14.7392 | -27.4791 | 23.7876 | 46.3795 | 17.7522 | 31.8945 | 23.2769 | -40.4497 | 49.3122 | -2.4879 | 50.2561 | -23.5702 | -0.15559 | -0.46839 | -0.16835 | -0.44681 | -0.2422 | -0.60005
| b (D65/2°) | 12.8495 | 17.1879 | -22.3149 | 22.3136 | -24.9077 | 1.1856 | 68.3122 | 58.8256 | -21.9629 | 14.8386 | -44.8458 | 57.0728 | -49.8628 | 32.6878 | 27.4783 | 81.234 | -16.0429 | -26.6585 | -0.83846 | -0.63475 | 0.43733 | 0.27004 | 0.50917 | 2.8198
Sensor
| L (D65/2°) | 38.5773 | 65.1317 | 49.8824 | 42.3596 | 54.9636 | 70.9463 | 71.568 | 72.5925 | 30.4605 | 50.5337 | 39.6897 | 61.3808 | 31.4878 | 54.5482 | 41.1689 | 81.9663 | 50.8044 | 50.9924 | 21.1842 | 35.8022 | 49.7592 | 66.6956 | 81.9845 | 95.6919
| a (D65/2°) | 10.7545 | 15.5735 | -0.089781 | -13.567 | 13.6831 | -32.7859 | 14.7755 | -28.2558 | 23.6759 | 46.0165 | 17.7232 | 32.25 | 22.0856 | -40.6269 | 49.3662 | -2.0232 | 50.6768 | -23.2142 | 0.22845 | -0.31706 | -0.021804 | -0.48741 | -0.3706 | -0.62379
| b (D65/2°) | 13.116 | 16.9731 | -21.5035 | 22 | -25.1413 | 1.6183 | 68.9404 | 58.2215 | -19.6638 | 14.4834 | -46.2277 | 58.6641 | -48.4997 | 31.9389 | 28.0201 | 82.5048 | -16.0676 | -26.4073 | -0.27863 | -0.47675 | 0.60662 | 0.37214 | 0.5207 | 2.4392
| Color | Dark Skin | Light Skin | Blue Sky | Foliage | Blue Flower | Bluish Green | Orange Yellow | Yellow Green | Purple | Moderate Red | Purplish Blue | Orange | Blue | Green | Red | Yellow | Magenta | Cyan | Black | Neutral 3.5 | Neutral 5 | Neutral 6.5 | Neutral 8 | White
| Delta E (D65/2°) | 0.79617 | 0.6034 | 1.1425 | 1.5019 | 0.56171 | 0.70667 | 0.69346 | 0.98554 | 2.3034 | 0.50792 | 1.9329 | 1.7176 | 2.1374 | 0.88224 | 0.65455 | 1.3787 | 0.48677 | 0.63225 | 0.69883 | 0.6053 | 1.3423 | 0.68031 | 0.195 | 0.4903
| Max DeltaE | 1: 2.3034
| Min Delta E | 1: 0.195
| Average DeltaE | 1: 0.98487