## Header
| Version | V1
| Published | 2021-04-06
Disclaimer - ams AG
This Excel sheet is a calibration aid that shows the calibration based on a spectral to XYZ calibration. All data's in Excel are examples and were measured based on an AS7352 EVK test board and GUI in combination with a LED based back light. The data's don't have a general validity for all sensor boards or only reduced accuracy for sensor boards of identical construction. In this case, accuracy can be increased by using a balancing (e.h. White). Algorithm for balancing are described in the app note "Methods for spectral calibration. Only use the original downloaded versions of this file and do not send, save, use, or manipulate a copy of it. Versions of this document with a different author or publication date are not the original file from ams and hence are not supported. All the tables are not read-only and therefore not tamper-proof. Thus every detail is only valid as information if it comes directly from the ams download and has not been modified.
IMPORTANT - PLEASE READ CAREFULLY BEFORE COPYING, INSTALLING, OR USING THE SOFTWARE. THIS SOFTWARE IS PROVIDED BY THE COPYRIGHT HOLDERS (ams AG – all rights reserved) AND CONTRIBUTORS "AS IS" AND ANY EXPRESS OR IMPLIED WARRANTIES, INCLUDING, BUT NOT LIMITED TO, THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE ARE DISCLAIMED. IN NO EVENT SHALL THE COPYRIGHT OWNER OR CONTRIBUTORS BE LIABLE FOR ANY DIRECT, INDIRECT, INCIDENTAL, SPECIAL, EXEMPLARY, OR CONSEQUENTIAL DAMAGES (INCLUDING, BUT NOT LIMITED TO PROCUREMENT OF SUBSTITUTE GOODS OR SERVICES; LOSS OF USE, DATA, OR PROFITS; OR BUSINESS INTERRUPTION) HOWEVER CAUSED AND ON ANY THEORY OF LIABILITY, WHETHER IN CONTRACT, STRICT LIABILITY, OR TORT (INCLUDING NEGLIGENCE OR OTHERWISE) ARISING IN ANY WAY OUT OF THE USE OF THIS SOFTWARE, EVEN IF ADVISED OF THE POSSIBILITY OF SUCH DAMAGE.
I) Slide Log_File inserts the measured sensor results (example log file) and data preparations for the CM calculations.
II) See the formulas in the sheet 'CM calculation' for the Calculation of a Correction Matrix based on the Spectrometer and Sensor results.
III) See the test results for colored targets, based on the actual calculated Correction Matrix, in the sheet 'Test Results', from line 50 and following.

## LogFile from Gui
[21 empty] | Extract from GUI protocol or log file | [6 empty] | Use this target to check our example or copy your own target in these colored cells.
Function | Comment | Sample num | Date | Time | AGAIN | Gain [x] | ATIME | ASTEP | TInt [ms] | Wait Time [ms] | LED Enable | LED Brightness [%] | LED Current [mA] | Measure Status | F1 (405nm) | F2 (424nm) | FZ (440nm) | F3 (473nm) | F4 (514nm) | FY (546nm) | F5 (546nm) | FXL (596nm) | F6 (635nm) | F7 (670nm) | F8 (710nm) | VIS | NIR | Basic F1 (405nm) | Basic F2 (424nm) | Basic FZ (440nm) | Basic F3 (473nm) | Basic F4 (514nm) | Basic FY (546nm) | Basic F5 (546nm) | Basic FXL (596nm) | Basic F6 (635nm) | Basic F7 (670nm) | Basic F8 (710nm) | Basic VIS | Basic NIR | Corr F1 (405nm) | Corr F2 (424nm) | Corr FZ (440nm) | Corr F3 (473nm) | Corr F4 (514nm) | Corr FY (546nm) | Corr F5 (546nm) | Corr FXL (596nm) | Corr F6 (635nm) | Corr F7 (670nm) | Corr F8 (710nm) | Corr VIS | Corr NIR | Temp Comp | Temp LED Bottom
Raw data |  | 1 | 06,04,2021 | 13:02 | 12 | 2048 | 11 | 61390 | 2048.004 | 0 | FALSE | 0 | 0 |  | 1315 | 2158 | 3780 | 3866 | 5279 | 5802 | 1794 | 5430 | 5597 | 5890 | 3884 | 0 | 0 | 0.000314 | 0.000515 | 0.000901 | 0.000922 | 0.001259 | 0.001383 | 0.000428 | 0.001295 | 0.001334 | 0.001404 | 0.000926 | 0 | 0 | 0.000331 | 0.000537 | 0.000928 | 0.000938 | 0.001264 | 0.001366 | 0.00041 | 0.001289 | 0.001354 | 0.001399 | 0.000864 | 0 | 0 | Off | 25.44
Raw data |  | 2 | 06,04,2021 | 13:02 | 12 | 2048 | 1 | 51326 | 285.378 | 0 | FALSE | 0 | 0 |  | 908 | 1635 | 2198 | 1960 | 2111 | 19090 | 867 | 39566 | 59163 | 23462 | 2565 | 0 | 0 | 0.001554 | 0.002797 | 0.003761 | 0.003354 | 0.003612 | 0.032663 | 0.001483 | 0.067697 | 0.10123 | 0.040143 | 0.004389 | 0 | 0 | 0.00164 | 0.002919 | 0.003872 | 0.003412 | 0.003628 | 0.03225 | 0.001421 | 0.067417 | 0.10271 | 0.040003 | 0.004095 | 0 | 0 | Off | 25.38
Raw data |  | 3 | 06,04,2021 | 13:03 | 12 | 2048 | 1 | 57930 | 322.096 | 0 | FALSE | 0 | 0 |  | 1081 | 2055 | 6353 | 6049 | 54250 | 59004 | 25042 | 37256 | 4319 | 2718 | 1256 | 0 | 0 | 0.001639 | 0.003115 | 0.009631 | 0.00917 | 0.08224 | 0.089447 | 0.037962 | 0.056478 | 0.006547 | 0.00412 | 0.001904 | 0 | 0 | 0.00173 | 0.003251 | 0.009916 | 0.00933 | 0.082604 | 0.088316 | 0.036353 | 0.056245 | 0.006643 | 0.004106 | 0.001777 | 0 | 0 | Off | 25.38
Raw data |  | 4 | 06,04,2021 | 13:03 | 12 | 2048 | 2 | 44438 | 370.621 | 0 | FALSE | 0 | 0 |  | 1163 | 14693 | 58986 | 30610 | 16914 | 9848 | 2127 | 3040 | 2003 | 1688 | 1031 | 0 | 0 | 0.001532 | 0.019358 | 0.077712 | 0.040328 | 0.022284 | 0.012974 | 0.002802 | 0.004005 | 0.002639 | 0.002224 | 0.001358 | 0 | 0 | 0.001617 | 0.0202 | 0.080011 | 0.041034 | 0.022382 | 0.01281 | 0.002683 | 0.003989 | 0.002677 | 0.002216 | 0.001267 | 0 | 0 | Off | 25.38
Raw data |  | 5 | 06,04,2021 | 13:03 | 12 | 2048 | 1 | 39319 | 218.619 | 0 | FALSE | 0 | 0 |  | 1838 | 11164 | 40889 | 24263 | 47775 | 58987 | 18431 | 55782 | 47793 | 19637 | 2519 | 0 | 0 | 0.004105 | 0.024935 | 0.091325 | 0.054191 | 0.1067 | 0.13175 | 0.041165 | 0.12459 | 0.10675 | 0.043859 | 0.005626 | 0 | 0 | 0.004333 | 0.026019 | 0.094026 | 0.05514 | 0.10718 | 0.13008 | 0.03942 | 0.12407 | 0.10831 | 0.043705 | 0.00525 | 0 | 0 | Off | 25.38
Raw data |  | 6 | 06,04,2021 | 13:03 | 12 | 2048 | 11 | 61390 | 2048.004 | 0 | FALSE | 0 | 0 |  | 1535 | 3520 | 8256 | 6217 | 9186 | 9812 | 3118 | 8635 | 8094 | 7024 | 4499 | 0 | 0 | 0.000366 | 0.000839 | 0.001968 | 0.001482 | 0.00219 | 0.002339 | 0.000743 | 0.002059 | 0.00193 | 0.001675 | 0.001073 | 0 | 0 | 0.000386 | 0.000876 | 0.002027 | 0.001508 | 0.0022 | 0.00231 | 0.000712 | 0.00205 | 0.001958 | 0.001669 | 0.001001 | 0 | 0 | Off | 25.38
Raw data |  | 7 | 06,04,2021 | 13:04 | 12 | 2048 | 1 | 56984 | 316.837 | 0 | FALSE | 0 | 0 |  | 1927 | 11564 | 41533 | 24309 | 47905 | 59004 | 18292 | 56196 | 49073 | 19993 | 2702 | 0 | 0 | 0.00297 | 0.017821 | 0.064007 | 0.037463 | 0.073827 | 0.090932 | 0.02819 | 0.086604 | 0.075627 | 0.030811 | 0.004164 | 0 | 0 | 0.003134 | 0.018597 | 0.0659 | 0.038119 | 0.074153 | 0.089782 | 0.026995 | 0.086246 | 0.076733 | 0.030704 | 0.003885 | 0 | 0 | Off | 25.38
Raw data |  | 8 | 06,04,2021 | 13:04 | 12 | 2048 | 3 | 64210 | 714.026 | 0 | FALSE | 0 | 0 |  | 2166 | 11454 | 40273 | 23285 | 47825 | 59173 | 18192 | 57075 | 50794 | 21250 | 3523 | 0 | 0 | 0.001481 | 0.007833 | 0.02754 | 0.015923 | 0.032705 | 0.040465 | 0.01244 | 0.03903 | 0.034735 | 0.014532 | 0.002409 | 0 | 0 | 0.001563 | 0.008174 | 0.028355 | 0.016202 | 0.032849 | 0.039953 | 0.011913 | 0.038869 | 0.035243 | 0.014481 | 0.002248 | 0 | 0 | Off | 25.38
Raw data |  | 9 | 06,04,2021 | 13:04 | 12 | 2048 | 11 | 61390 | 2048.004 | 0 | FALSE | 0 | 0 |  | 3037 | 13808 | 46085 | 26972 | 48982 | 57715 | 17879 | 54094 | 48153 | 22495 | 6124 | 0 | 0 | 0.000724 | 0.003292 | 0.010988 | 0.006431 | 0.011678 | 0.01376 | 0.004263 | 0.012897 | 0.011481 | 0.005363 | 0.00146 | 0 | 0 | 0.000764 | 0.003435 | 0.011312 | 0.006543 | 0.01173 | 0.013586 | 0.004082 | 0.012844 | 0.011648 | 0.005344 | 0.001362 | 0 | 0 | Off | 25.38
Raw data |  | 10 | 06,04,2021 | 13:04 | 12 | 2048 | 11 | 61390 | 2048.004 | 0 | FALSE | 0 | 0 |  | 2068 | 7153 | 21510 | 13491 | 22611 | 25839 | 8007 | 23932 | 21903 | 12314 | 5054 | 0 | 0 | 0.000493 | 0.001705 | 0.005128 | 0.003216 | 0.005391 | 0.00616 | 0.001909 | 0.005706 | 0.005222 | 0.002936 | 0.001205 | 0 | 0 | 0.00052 | 0.00178 | 0.00528 | 0.003273 | 0.005415 | 0.006083 | 0.001828 | 0.005682 | 0.005298 | 0.002926 | 0.001124 | 0 | 0 | Off | 25.38
Raw data |  | 11 | 06,04,2021 | 13:04 | 12 | 2048 | 1 | 42484 | 236.217 | 0 | FALSE | 0 | 0 |  | 1398 | 2478 | 5591 | 5709 | 41537 | 58361 | 18776 | 59062 | 50938 | 20808 | 2561 | 0 | 0 | 0.00289 | 0.005122 | 0.011557 | 0.011801 | 0.085861 | 0.12064 | 0.038812 | 0.12209 | 0.10529 | 0.043012 | 0.005294 | 0 | 0 | 0.00305 | 0.005345 | 0.011899 | 0.012008 | 0.08624 | 0.11911 | 0.037166 | 0.12158 | 0.10683 | 0.042862 | 0.00494 | 0 | 0 | Off | 25.38
Raw data |  | 12 | 06,04,2021 | 13:04 | 12 | 2048 | 1 | 51212 | 284.744 | 0 | FALSE | 0 | 0 |  | 1606 | 12728 | 48131 | 25597 | 14071 | 25200 | 2085 | 40202 | 59019 | 23750 | 2747 | 0 | 0 | 0.002754 | 0.021826 | 0.082535 | 0.043894 | 0.024129 | 0.043213 | 0.003575 | 0.068939 | 0.10121 | 0.040727 | 0.004711 | 0 | 0 | 0.002907 | 0.022776 | 0.084976 | 0.044662 | 0.024236 | 0.042667 | 0.003424 | 0.068653 | 0.10269 | 0.040584 | 0.004395 | 0 | 0 | Off | 25.44
Raw data |  | 13 | 06,04,2021 | 13:04 | 12 | 2048 | 1 | 50613 | 281.414 | 0 | FALSE | 0 | 0 |  | 1641 | 12897 | 50521 | 28222 | 58917 | 57795 | 23169 | 33695 | 4237 | 2733 | 1273 | 0 | 0 | 0.002847 | 0.022378 | 0.087659 | 0.048968 | 0.10223 | 0.10028 | 0.040201 | 0.058464 | 0.007352 | 0.004742 | 0.002209 | 0 | 0 | 0.003005 | 0.023351 | 0.090252 | 0.049825 | 0.10268 | 0.099012 | 0.038496 | 0.058222 | 0.007459 | 0.004725 | 0.002061 | 0 | 0 | Off | 25.44
[27 empty] | Sensor Data's | Dark | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan
[27 empty] | Basic F1 (405nm) | 0.000314 | 0.001554 | 0.001639 | 0.001532 | 0.004105 | 0.000366 | 0.00297 | 0.001481 | 0.000724 | 0.000493 | 0.00289 | 0.002754 | 0.002847
[27 empty] | Basic F2 (424nm) | 0.000515 | 0.002797 | 0.003115 | 0.019358 | 0.024935 | 0.000839 | 0.017821 | 0.007833 | 0.003292 | 0.001705 | 0.005122 | 0.021826 | 0.022378
[27 empty] | Basic FZ (440nm) | 0.000901 | 0.003761 | 0.009631 | 0.077712 | 0.091325 | 0.001968 | 0.064007 | 0.02754 | 0.010988 | 0.005128 | 0.011557 | 0.082535 | 0.087659
[27 empty] | Basic F3 (473nm) | 0.000922 | 0.003354 | 0.00917 | 0.040328 | 0.054191 | 0.001482 | 0.037463 | 0.015923 | 0.006431 | 0.003216 | 0.011801 | 0.043894 | 0.048968
[27 empty] | Basic F4 (514nm) | 0.001259 | 0.003612 | 0.08224 | 0.022284 | 0.1067 | 0.00219 | 0.073827 | 0.032705 | 0.011678 | 0.005391 | 0.085861 | 0.024129 | 0.10223
[27 empty] | Basic FY (546nm) | 0.001383 | 0.032663 | 0.089447 | 0.012974 | 0.13175 | 0.002339 | 0.090932 | 0.040465 | 0.01376 | 0.00616 | 0.12064 | 0.043213 | 0.10028
[27 empty] | Basic F5 (546nm) | 0.000428 | 0.001483 | 0.037962 | 0.002802 | 0.041165 | 0.000743 | 0.02819 | 0.01244 | 0.004263 | 0.001909 | 0.038812 | 0.003575 | 0.040201
[27 empty] | Basic FXL (596nm) | 0.001295 | 0.067697 | 0.056478 | 0.004005 | 0.12459 | 0.002059 | 0.086604 | 0.03903 | 0.012897 | 0.005706 | 0.12209 | 0.068939 | 0.058464
[27 empty] | Basic F6 (635nm) | 0.001334 | 0.10123 | 0.006547 | 0.002639 | 0.10675 | 0.00193 | 0.075627 | 0.034735 | 0.011481 | 0.005222 | 0.10529 | 0.10121 | 0.007352
[27 empty] | Basic F7 (670nm) | 0.001404 | 0.040143 | 0.00412 | 0.002224 | 0.043859 | 0.001675 | 0.030811 | 0.014532 | 0.005363 | 0.002936 | 0.043012 | 0.040727 | 0.004742
[27 empty] | Basic F8 (710nm) | 0.000926 | 0.004389 | 0.001904 | 0.001358 | 0.005626 | 0.001073 | 0.004164 | 0.002409 | 0.00146 | 0.001205 | 0.005294 | 0.004711 | 0.002209
[27 empty] | Basic VIS | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
[27 empty] | Basic NIR | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## CM Calculation
Calculation 'Correction Matrix' |  | Yellow cells = input cells
TARGET DATA T (Reference Values) | Reference values | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
|  | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan
| X | 0.29967 | 0.33962 | 0.15326 | 0.68109 | 0.015069 | 0.48163 | 0.22162 | 0.080848 | 0.038389 | 0.54283 | 0.4364 | 0.39028
| Y | 0.16242 | 0.49082 | 0.072889 | 0.70119 | 0.01585 | 0.48982 | 0.22457 | 0.080976 | 0.038495 | 0.64025 | 0.22028 | 0.5506
| Z | 0.022659 | 0.069815 | 0.7529 | 0.78105 | 0.022629 | 0.56212 | 0.25117 | 0.10793 | 0.053676 | 0.073108 | 0.72606 | 0.78049
Offset Sensor (e.g. Optical Overcrossing) | Offset | Yellow cells = input cells 1: Sensor; Yellow cells = input cells 3: Background of the calibration is a calculation of target values T and sensor values S with the method of linear regression. The target values are in XYZ, measured with a spectrometer or known standard values of colors, which are measured with a sensor as sensor raw data. The sensor raw data are corrected (only offset is used here) and then referenced with the spectrometer data. The result of the calibration is a correction matrix CM, which can be used under the same conditions as in the calibration matching of the sensor values. Realize action 1...4 in the calculation of a customized calibration matrix or follow the formulas backwards from calibration matrix to input values to understand the calculations.
| 400 | Yellow cells = input cells 1: 0.000314
| 424 | Yellow cells = input cells 1: 0.000515
| 450 | Yellow cells = input cells 1: 0.000901
| 473 | Yellow cells = input cells 1: 0.000922
| 514 | Yellow cells = input cells 1: 0.001259
| 555 | Yellow cells = input cells 1: 0.001383
| 595 | Yellow cells = input cells 1: 0.001295
| 635 | Yellow cells = input cells 1: 0.001334
| 685 | Yellow cells = input cells 1: 0.001404
| 745 | Yellow cells = input cells 1: 0.000926
Sensor Matrix s | Basic Counts | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
|  | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan
| 400 | 0.001554 | 0.001639 | 0.001532 | 0.004105 | 0.000366 | 0.00297 | 0.001481 | 0.000724 | 0.000493 | 0.00289 | 0.002754 | 0.002847
| 424 | 0.002797 | 0.003115 | 0.019358 | 0.024935 | 0.000839 | 0.017821 | 0.007833 | 0.003292 | 0.001705 | 0.005122 | 0.021826 | 0.022378
| 450 | 0.003761 | 0.009631 | 0.077712 | 0.091325 | 0.001968 | 0.064007 | 0.02754 | 0.010988 | 0.005128 | 0.011557 | 0.082535 | 0.087659
| 473 | 0.003354 | 0.00917 | 0.040328 | 0.054191 | 0.001482 | 0.037463 | 0.015923 | 0.006431 | 0.003216 | 0.011801 | 0.043894 | 0.048968
| 514 | 0.003612 | 0.08224 | 0.022284 | 0.1067 | 0.00219 | 0.073827 | 0.032705 | 0.011678 | 0.005391 | 0.085861 | 0.024129 | 0.10223
| 555 | 0.032663 | 0.089447 | 0.012974 | 0.13175 | 0.002339 | 0.090932 | 0.040465 | 0.01376 | 0.00616 | 0.12064 | 0.043213 | 0.10028
| 595 | 0.067697 | 0.056478 | 0.004005 | 0.12459 | 0.002059 | 0.086604 | 0.03903 | 0.012897 | 0.005706 | 0.12209 | 0.068939 | 0.058464
| 635 | 0.10123 | 0.006547 | 0.002639 | 0.10675 | 0.00193 | 0.075627 | 0.034735 | 0.011481 | 0.005222 | 0.10529 | 0.10121 | 0.007352
| 685 | 0.040143 | 0.00412 | 0.002224 | 0.043859 | 0.001675 | 0.030811 | 0.014532 | 0.005363 | 0.002936 | 0.043012 | 0.040727 | 0.004742
| 745 | 0.004389 | 0.001904 | 0.001358 | 0.005626 | 0.001073 | 0.004164 | 0.002409 | 0.00146 | 0.001205 | 0.005294 | 0.004711 | 0.002209
Sensor Matrix corr offset S' = S - offset | Basic-Offset | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan
| 400 | 0.00124 | 0.001325 | 0.001218 | 0.003791 | 5.2e-05 | 0.002656 | 0.001167 | 0.00041 | 0.000179 | 0.002576 | 0.00244 | 0.002533
| 424 | 0.002282 | 0.0026 | 0.018843 | 0.02442 | 0.000324 | 0.017306 | 0.007318 | 0.002777 | 0.00119 | 0.004607 | 0.021311 | 0.021863
| 450 | 0.00286 | 0.00873 | 0.076811 | 0.090424 | 0.001067 | 0.063106 | 0.026639 | 0.010087 | 0.004227 | 0.010656 | 0.081634 | 0.086758
| 473 | 0.002432 | 0.008248 | 0.039406 | 0.053269 | 0.00056 | 0.036541 | 0.015001 | 0.005509 | 0.002294 | 0.010879 | 0.042972 | 0.048046
| 514 | 0.002353 | 0.080981 | 0.021025 | 0.10544 | 0.000931 | 0.072568 | 0.031446 | 0.010419 | 0.004132 | 0.084602 | 0.02287 | 0.10097
| 555 | 0.03128 | 0.088064 | 0.011591 | 0.13036 | 0.000956 | 0.089549 | 0.039082 | 0.012377 | 0.004777 | 0.11925 | 0.04183 | 0.098897
| 595 | 0.066402 | 0.055183 | 0.00271 | 0.12329 | 0.000764 | 0.085309 | 0.037735 | 0.011602 | 0.004411 | 0.12079 | 0.067644 | 0.057169
| 635 | 0.099894 | 0.005213 | 0.001305 | 0.10541 | 0.000596 | 0.074293 | 0.033401 | 0.010147 | 0.003888 | 0.10396 | 0.099872 | 0.006018
| 685 | 0.038739 | 0.002716 | 0.00082 | 0.042455 | 0.000271 | 0.029407 | 0.013128 | 0.003959 | 0.001532 | 0.041608 | 0.039323 | 0.003338
| 745 | 0.003463 | 0.000978 | 0.000432 | 0.0047 | 0.000147 | 0.003238 | 0.001483 | 0.000534 | 0.000279 | 0.004368 | 0.003785 | 0.001283
|  | F1 | F2 | FZ | F3 | F4 | FY | FXL | F6 | F7 | F8
Calculation: Matrix for linear transformation CM |  | 400 | 424 | 450 | 473 | 514 | 553 | 595 | 635 | 685 | 754
| A = T * S'_trans | 0.0086206 | 0.051653 | 0.18974 | 0.11073 | 0.24149 | 0.30442 | 0.291 | 0.25023 | 0.10008 | 0.011105
|  | 0.0087841 | 0.050221 | 0.18422 | 0.10936 | 0.27402 | 0.33442 | 0.29962 | 0.22954 | 0.092329 | 0.010682
|  | 0.0097766 | 0.078304 | 0.3004 | 0.16644 | 0.25555 | 0.29545 | 0.2652 | 0.22222 | 0.089386 | 0.01048
| B = [S' * S'_trans]^-1 | 43427666415.0858 | -3331533091.69 | 48089964.1475 | 852418712.7461 | -664757494.9415 | -1267019152.6746 | 2227985575.4293 | -708219797.9676 | -1662951285.2043 | -5786281360.2963
|  | -3331533091.6175 | 281047177.3057 | -7127545.9566 | -70108862.586 | 48726023.8004 | 98251083.2228 | -169224431.4043 | 46343115.3437 | 144238868.7624 | 437081502.5136
|  | 48089964.2448 | -7127545.9644 | 1688567.0687 | -124388.7762 | 1806911.5154 | -9253500.6159 | 11718725.3537 | -4039306.4118 | -2826107.7624 | -8007047.8662
|  | 852418712.5899 | -70108862.575 | -124388.7784 | 20265103.459 | -15875446.9994 | -14741786.9084 | 31074727.1616 | -6676996.6715 | -37957809.1703 | -107964049.2475
|  | -664757494.724 | 48726023.7842 | 1806911.5172 | -15875446.9974 | 15385489.3589 | 2888875.2578 | -14489422.8789 | 2712684.328 | 26242458.2798 | 85114484.9416
|  | -1267019153.4858 | 98251083.2896 | -9253500.6142 | -14741786.9297 | 2888875.2766 | 96020398.1201 | -138314414.7913 | 52979185.3021 | 43065718.0974 | 179168883.1015
|  | 2227985576.4502 | -169224431.4894 | 11718725.3502 | 31074727.1906 | -14489422.9057 | -138314414.7795 | 206611489.4383 | -77812152.6332 | -77357671.9163 | -309406646.2425
|  | -708219798.448 | 46343115.3831 | -4039306.411 | -6676996.6837 | 2712684.3387 | 52979185.3038 | -77812152.6423 | 32752110.8071 | 17472966.276 | 102052726.8543
|  | -1662951285.0594 | 144238868.7538 | -2826107.7582 | -37957809.1738 | 26242458.2865 | 43065718.0598 | -77357671.8668 | 17472966.254 | 79530813.816 | 212172642.7007
|  | -5786281360.4628 | 437081502.5364 | -8007047.8537 | -107964049.2714 | 85114484.9728 | 179168882.9996 | -309406646.1167 | 102052726.7936 | 212172642.7267 | 805759351.9487
Calculated Correction Matrix | CM = A * B | 65.0174 | -301.1651 | 42.9979 | 57.691 | 21.8783 | -32.9491 | 22.9521 | 90.1675 | -254.6547 | 275.4698
|  | 273.9534 | -18.9919 | 0.54285 | 2.3928 | 0.87803 | -8.4192 | 15.3079 | 1.8134 | -33.2361 | 60.9499
|  | -14.5036 | 5.5535 | 32.8973 | -43.0367 | 58.696 | -160.0867 | 176.3384 | -74.5359 | 4.6274 | 132.4464
|  | 400 | 424 | 450 | 473 | 514 | 553 | 546 | 595 | 635 | 685 | 754 | 750
CM File | Wavelength | F1 | F2 | FZ | F3 | F4 | FY | F5 | FXL | F6 | F7 | F8 | NIR | VIS | Copy this matrix CM into the CM - File of AS7352 EVK when using a customized calibration.
| X | 65.0174 | -301.1651 | 42.9979 | 57.691 | 21.8783 | -32.9491 | 0 | 22.9521 | 90.1675 | -254.6547 | 275.4698 | 0 | 0
| Y | 273.9534 | -18.9919 | 0.54285 | 2.3928 | 0.87803 | -8.4192 | 0 | 15.3079 | 1.8134 | -33.2361 | 60.9499 | 0 | 0
| Z | -14.5036 | 5.5535 | 32.8973 | -43.0367 | 58.696 | -160.0867 | 0 | 176.3384 | -74.5359 | 4.6274 | 132.4464 | 0 | 0
Mask File | Title | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan |  |  | Copy this matrix into MaskFile of AS7352 EVK when using a customized calibration.
| X | 0.29967 | 0.33962 | 0.15326 | 0.68109 | 0.015069 | 0.48163 | 0.22162 | 0.080848 | 0.038389 | 0.54283 | 0.4364 | 0.39028
| Y | 0.16242 | 0.49082 | 0.072889 | 0.70119 | 0.01585 | 0.48982 | 0.22457 | 0.080976 | 0.038495 | 0.64025 | 0.22028 | 0.5506
| Z | 0.022659 | 0.069815 | 0.7529 | 0.78105 | 0.022629 | 0.56212 | 0.25117 | 0.10793 | 0.053676 | 0.073108 | 0.72606 | 0.78049
0.012238

## Test Results
Yellow cells = input cells | Calculation 'Test Results'
Spectrometer | Reference values | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Spectrometer | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| X | 0.29967 | 0.33962 | 0.15326 | 0.68109 | 0.015069 | 0.48163 | 0.22162 | 0.080848 | 0.038389 | 0.54283 | 0.4364 | 0.39028 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Y | 0.16242 | 0.49082 | 0.072889 | 0.70119 | 0.01585 | 0.48982 | 0.22457 | 0.080976 | 0.038495 | 0.64025 | 0.22028 | 0.5506 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Z | 0.022659 | 0.069815 | 0.7529 | 0.78105 | 0.022629 | 0.56212 | 0.25117 | 0.10793 | 0.053676 | 0.073108 | 0.72606 | 0.78049 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| x | 0.6182 | 0.37725 | 0.15654 | 0.31483 | 0.28141 | 0.31406 | 0.3178 | 0.29971 | 0.29403 | 0.43212 | 0.31561 | 0.22673 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| y | 0.33506 | 0.5452 | 0.074449 | 0.32413 | 0.296 | 0.3194 | 0.32203 | 0.30018 | 0.29485 | 0.50968 | 0.15931 | 0.31986 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| z | 0.046744 | 0.07755 | 0.76901 | 0.36104 | 0.42259 | 0.36654 | 0.36017 | 0.40011 | 0.41112 | 0.058198 | 0.52509 | 0.45341 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Sensor Values | [7 empty] | The correction matrix is used to match the immediate sensor results in the application - here the color space is used to define target colors and to check the accuracy with reference values. This can be done with the colors and values from the calibration or with other colors, each measured with spectrometer and sensor (follow action 1...3). The data chain in the sheet is that the sensor values multiplied by the matrix results in the sensor XYZ values. These are compared with the spectrometer values as XYZ or further down as xy and Lab. The errors are output as Delta per color or Max/Min/Average.
| Offset | Dark Basic Counts
| 400 | 0.000314
| 424 | 0.000515
| 450 | 0.000901
| 473 | 0.000922
| 514 | 0.001259
| 555 | 0.001383
| 595 | 0.001295
| 635 | 0.001334
| 685 | 0.001404
| 745 | 0.000926
| Sensor Values | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
| Basic counts | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 400 | 0.001554 | 0.001639 | 0.001532 | 0.004105 | 0.000366 | 0.00297 | 0.001481 | 0.000724 | 0.000493 | 0.00289 | 0.002754 | 0.002847 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 424 | 0.002797 | 0.003115 | 0.019358 | 0.024935 | 0.000839 | 0.017821 | 0.007833 | 0.003292 | 0.001705 | 0.005122 | 0.021826 | 0.022378 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 450 | 0.003761 | 0.009631 | 0.077712 | 0.091325 | 0.001968 | 0.064007 | 0.02754 | 0.010988 | 0.005128 | 0.011557 | 0.082535 | 0.087659 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 473 | 0.003354 | 0.00917 | 0.040328 | 0.054191 | 0.001482 | 0.037463 | 0.015923 | 0.006431 | 0.003216 | 0.011801 | 0.043894 | 0.048968 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 514 | 0.003612 | 0.08224 | 0.022284 | 0.1067 | 0.00219 | 0.073827 | 0.032705 | 0.011678 | 0.005391 | 0.085861 | 0.024129 | 0.10223 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 555 | 0.032663 | 0.089447 | 0.012974 | 0.13175 | 0.002339 | 0.090932 | 0.040465 | 0.01376 | 0.00616 | 0.12064 | 0.043213 | 0.10028 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 595 | 0.067697 | 0.056478 | 0.004005 | 0.12459 | 0.002059 | 0.086604 | 0.03903 | 0.012897 | 0.005706 | 0.12209 | 0.068939 | 0.058464 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 635 | 0.10123 | 0.006547 | 0.002639 | 0.10675 | 0.00193 | 0.075627 | 0.034735 | 0.011481 | 0.005222 | 0.10529 | 0.10121 | 0.007352 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 685 | 0.040143 | 0.00412 | 0.002224 | 0.043859 | 0.001675 | 0.030811 | 0.014532 | 0.005363 | 0.002936 | 0.043012 | 0.040727 | 0.004742 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 745 | 0.004389 | 0.001904 | 0.001358 | 0.005626 | 0.001073 | 0.004164 | 0.002409 | 0.00146 | 0.001205 | 0.005294 | 0.004711 | 0.002209 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Corrected counts | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 400 | 0.00124 | 0.001325 | 0.001218 | 0.003791 | 5.2e-05 | 0.002656 | 0.001167 | 0.00041 | 0.000179 | 0.002576 | 0.00244 | 0.002533 | -0.000314 | -0.000314 | -0.000314 | -0.000314 | -0.000314 | -0.000314 | -0.000314 | -0.000314 | -0.000314 | -0.000314 | -0.000314 | -0.000314
| 424 | 0.002282 | 0.0026 | 0.018843 | 0.02442 | 0.000324 | 0.017306 | 0.007318 | 0.002777 | 0.00119 | 0.004607 | 0.021311 | 0.021863 | -0.000515 | -0.000515 | -0.000515 | -0.000515 | -0.000515 | -0.000515 | -0.000515 | -0.000515 | -0.000515 | -0.000515 | -0.000515 | -0.000515
| 450 | 0.00286 | 0.00873 | 0.076811 | 0.090424 | 0.001067 | 0.063106 | 0.026639 | 0.010087 | 0.004227 | 0.010656 | 0.081634 | 0.086758 | -0.000901 | -0.000901 | -0.000901 | -0.000901 | -0.000901 | -0.000901 | -0.000901 | -0.000901 | -0.000901 | -0.000901 | -0.000901 | -0.000901
| 473 | 0.002432 | 0.008248 | 0.039406 | 0.053269 | 0.00056 | 0.036541 | 0.015001 | 0.005509 | 0.002294 | 0.010879 | 0.042972 | 0.048046 | -0.000922 | -0.000922 | -0.000922 | -0.000922 | -0.000922 | -0.000922 | -0.000922 | -0.000922 | -0.000922 | -0.000922 | -0.000922 | -0.000922
| 514 | 0.002353 | 0.080981 | 0.021025 | 0.10544 | 0.000931 | 0.072568 | 0.031446 | 0.010419 | 0.004132 | 0.084602 | 0.02287 | 0.10097 | -0.001259 | -0.001259 | -0.001259 | -0.001259 | -0.001259 | -0.001259 | -0.001259 | -0.001259 | -0.001259 | -0.001259 | -0.001259 | -0.001259
| 555 | 0.03128 | 0.088064 | 0.011591 | 0.13036 | 0.000956 | 0.089549 | 0.039082 | 0.012377 | 0.004777 | 0.11925 | 0.04183 | 0.098897 | -0.001383 | -0.001383 | -0.001383 | -0.001383 | -0.001383 | -0.001383 | -0.001383 | -0.001383 | -0.001383 | -0.001383 | -0.001383 | -0.001383
| 595 | 0.066402 | 0.055183 | 0.00271 | 0.12329 | 0.000764 | 0.085309 | 0.037735 | 0.011602 | 0.004411 | 0.12079 | 0.067644 | 0.057169 | -0.001295 | -0.001295 | -0.001295 | -0.001295 | -0.001295 | -0.001295 | -0.001295 | -0.001295 | -0.001295 | -0.001295 | -0.001295 | -0.001295
| 635 | 0.099894 | 0.005213 | 0.001305 | 0.10541 | 0.000596 | 0.074293 | 0.033401 | 0.010147 | 0.003888 | 0.10396 | 0.099872 | 0.006018 | -0.001334 | -0.001334 | -0.001334 | -0.001334 | -0.001334 | -0.001334 | -0.001334 | -0.001334 | -0.001334 | -0.001334 | -0.001334 | -0.001334
| 685 | 0.038739 | 0.002716 | 0.00082 | 0.042455 | 0.000271 | 0.029407 | 0.013128 | 0.003959 | 0.001532 | 0.041608 | 0.039323 | 0.003338 | -0.001404 | -0.001404 | -0.001404 | -0.001404 | -0.001404 | -0.001404 | -0.001404 | -0.001404 | -0.001404 | -0.001404 | -0.001404 | -0.001404
| 745 | 0.003463 | 0.000978 | 0.000432 | 0.0047 | 0.000147 | 0.003238 | 0.001483 | 0.000534 | 0.000279 | 0.004368 | 0.003785 | 0.001283 | -0.000926 | -0.000926 | -0.000926 | -0.000926 | -0.000926 | -0.000926 | -0.000926 | -0.000926 | -0.000926 | -0.000926 | -0.000926 | -0.000926
Sensor Values and Results | Sensor Corrected Results | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| X | 0.29761 | 0.3388 | 0.14856 | 0.68266 | 0.015616 | 0.47952 | 0.22626 | 0.082144 | 0.038889 | 0.54116 | 0.43845 | 0.39117 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| Y | 0.16361 | 0.49129 | 0.075599 | 0.70029 | 0.015509 | 0.49103 | 0.22193 | 0.08015 | 0.038291 | 0.64121 | 0.2191 | 0.55009 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| Z | 0.016164 | 0.067259 | 0.73808 | 0.78597 | 0.024672 | 0.55554 | 0.26544 | 0.11299 | 0.054216 | 0.06789 | 0.73253 | 0.78327 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| x | 0.62343 | 0.37756 | 0.15439 | 0.31475 | 0.27988 | 0.31422 | 0.31706 | 0.2984 | 0.29597 | 0.43284 | 0.31541 | 0.22683 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| y | 0.34271 | 0.54749 | 0.078565 | 0.32288 | 0.27795 | 0.32176 | 0.31098 | 0.29115 | 0.29142 | 0.51286 | 0.15762 | 0.31898 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| z | 0.033859 | 0.074953 | 0.76705 | 0.36238 | 0.44217 | 0.36403 | 0.37196 | 0.41045 | 0.41261 | 0.054301 | 0.52697 | 0.45419 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Diff XYZ | XSP-XS | 0.6859% | 0.24% | 3.069% | -0.2302% | -3.633% | 0.4377% | -2.094% | -1.603% | -1.301% | 0.3067% | -0.469% | -0.227% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
| YSP-XS | -0.73% | -0.09547% | -3.718% | 0.1286% | 2.153% | -0.2469% | 1.178% | 1.02% | 0.529% | -0.1495% | 0.5361% | 0.09251% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
| ZSP-XS | 28.67% | 3.661% | 1.968% | -0.6298% | -9.028% | 1.171% | -5.681% | -4.688% | -1.005% | 7.137% | -0.8914% | -0.3559% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
|  | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Diff xy | dxy | 0.0092706 | 0.0023068 | 0.0046462 | 0.0012534 | 0.018112 | 0.0023631 | 0.011071 | 0.0091257 | 0.003932 | 0.0032608 | 0.0017015 | 0.00088608 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Average dxy | 0.000926 1: 0.0028304
| Max dxy | 0.000926 1: 0.018112
Lab results | Spectrometer
| L | 1.4671 | 4.4336 | 0.6584 | 6.3338 | 0.14317 | 4.4245 | 2.0285 | 0.73145 | 0.34772 | 5.7834 | 1.9898 | 4.9735 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| u' | 0.4275 | 0.17171 | 0.17489 | 0.20118 | 0.18795 | 0.20247 | 0.20409 | 0.19971 | 0.19767 | 0.20947 | 0.29493 | 0.14204 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| v' | 0.52133 | 0.55836 | 0.18715 | 0.46601 | 0.4448 | 0.46329 | 0.4653 | 0.45007 | 0.44598 | 0.55588 | 0.33495 | 0.45087 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Sensor
| L | 1.4778 | 4.4378 | 0.68288 | 6.3257 | 0.14009 | 4.4355 | 2.0046 | 0.72399 | 0.34588 | 5.792 | 1.9791 | 4.9689 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| u' | 0.42513 | 0.17133 | 0.16993 | 0.2016 | 0.19383 | 0.20166 | 0.20799 | 0.2024 | 0.20048 | 0.20888 | 0.29612 | 0.14234 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| v' | 0.52584 | 0.55899 | 0.19458 | 0.46531 | 0.43312 | 0.46462 | 0.459 | 0.44435 | 0.44415 | 0.55688 | 0.33295 | 0.45039 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| Deviation | red | green | blue | white | black | grey 25 | grey 50 | grey 75 | grey 88 | yellow | magenta | cyan | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Delta u'v' | 0.000926 1: 0.0050916; 0.000926 2: 0.00074104; 0.000926 3: 0.0089306; 0.000926 4: 0.00081312; 0.000926 5: 0.013079; 0.000926 6: 0.0015517; 0.000926 7: 0.0074115; 0.000926 8: 0.0063141; 0.000926 9: 0.0033532; 0.000926 10: 0.0011489; 0.000926 11: 0.0023358; 0.000926 12: 0.00056868
| Delta E | 0.000926 1: 0.011859; 0.000926 2: 0.0042973; 0.000926 3: 0.026055; 0.000926 4: 0.0081842; 0.000926 5: 0.013438; 0.000926 6: 0.011035; 0.000926 7: 0.025015; 0.000926 8: 0.0097749; 0.000926 9: 0.0038246; 0.000926 10: 0.0087203; 0.000926 11: 0.010921; 0.000926 12: 0.0046362
| Max u'v' | 0.000926 1: 0.013079
| Average Delta u'v' | 0.000926 1: 0.0042783
| Max DeltaE | 0.000926 1: 0.026055
| Average DeltaE | 0.000926 1: 0.01148